## Info
Sheets | Description | Comments
Configuration | Metering station configuration
ProcessData | Flow rates, Temperature, Pressure vs time
Composition | Fluid composition vs time
USM_unc | Input data for USM uncertainty analysis
T_unc | Input data for temperature transmitter uncertainty analysis
P_unc | Input data for pressure transmitter uncertainty analysis
Composition_unc | Uncertainty data for composition | Detailed uncertainties should be calculated by SW based on more general input (e.g. standard and representability)

## Configuration
Configuration | Setting | Unit | Comment
Quantity_meter | USM
Meter_configuration | Single
Quality_measurement | Online GC
Sampling | NA
Performance_requirements | OIML R 140
EoS | GERG-2008
Pipe_diameter | 9 | inch
Base_conditions | ISO13443
Note on A2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: quantity meter? Reply: Thanks. Text is now updated.
Note on C8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I think this should wither be a drop-down menu with some specified options or a single pre-determined unit. If we allow the user to insert anything here, it can easily get messy when handling all the options

## ProcessData
Case | DateTime | quantity_m3 | T_C | P_Mpa | Uptime_h
1 | 2023-03-06 00:15 | 795.78 | 7.2716 | 53.5637 | 1
2 | 2023-03-06 00:30 | 795.78 | 7.2898 | 53.5792 | 1
3 | 2023-03-06 00:45 | 796.78 | 7.2901 | 53.5997 | 1
4 | 2023-03-06 01:00 | 795.78 | 7.2752 | 53.5945 | 1
5 | 2023-03-06 01:15 | 793.78 | 7.2041 | 53.4695 | 1
6 | 2023-03-06 01:30 | 795.78 | 7.1899 | 53.3628 | 1
7 | 2023-03-06 01:45 | 798.77 | 7.195 | 53.3576 | 1
8 | 2023-03-06 02:00 | 797.78 | 7.2056 | 53.3708 | 1
9 | 2023-03-06 02:15 | 797.77 | 7.2082 | 53.3771 | 1
10 | 2023-03-06 02:30 | 796.78 | 7.2438 | 53.3879 | 1
11 | 2023-03-06 02:45 | 797.78 | 7.2385 | 53.4016 | 1
12 | 2023-03-06 03:00 | 796.77 | 7.2395 | 53.3891 | 1
13 | 2023-03-06 03:15 | 796.78 | 7.2236 | 53.3804 | 1
14 | 2023-03-06 03:30 | 797.78 | 7.2193 | 53.3687 | 1
15 | 2023-03-06 03:45 | 799.77 | 7.2157 | 53.363 | 1
16 | 2023-03-06 04:00 | 800.77 | 7.2563 | 53.3873 | 1
17 | 2023-03-06 04:15 | 801.76 | 7.2344 | 53.3747 | 1
18 | 2023-03-06 04:30 | 802.77 | 7.2337 | 53.3628 | 1
19 | 2023-03-06 04:45 | 802.76 | 7.2554 | 53.3844 | 1
20 | 2023-03-06 05:00 | 802.77 | 7.2542 | 53.4142 | 1
21 | 2023-03-06 05:15 | 801.76 | 7.2542 | 53.4162 | 1
22 | 2023-03-06 05:30 | 801.77 | 7.2258 | 53.3681 | 1
23 | 2023-03-06 05:45 | 802.76 | 7.2492 | 53.384 | 1
24 | 2023-03-06 06:00 | 802.77 | 7.2628 | 53.3933 | 1
25 | 2023-03-06 06:15 | 801.77 | 7.2511 | 53.38 | 1
26 | 2023-03-06 06:30 | 802.76 | 7.2524 | 53.3737 | 1
27 | 2023-03-06 06:45 | 803.76 | 7.2542 | 53.3927 | 1
28 | 2023-03-06 07:00 | 801.77 | 7.2567 | 53.3808 | 1
29 | 2023-03-06 07:15 | 803.76 | 7.2447 | 53.4119 | 1
30 | 2023-03-06 07:30 | 802.77 | 7.2361 | 53.3627 | 1
31 | 2023-03-06 07:45 | 801.76 | 7.2271 | 53.3695 | 1
32 | 2023-03-06 08:00 | 798.77 | 7.252 | 53.3673 | 1
33 | 2023-03-06 08:15 | 799.77 | 7.237 | 53.3621 | 1
34 | 2023-03-06 08:30 | 799.78 | 7.2417 | 53.3454 | 1
35 | 2023-03-06 08:45 | 799.77 | 7.2456 | 53.3942 | 1
36 | 2023-03-06 09:00 | 803.76 | 7.283 | 53.4012 | 1
37 | 2023-03-06 09:15 | 803.76 | 7.2623 | 53.3939 | 1
38 | 2023-03-06 09:30 | 803.76 | 7.2671 | 53.397 | 1
39 | 2023-03-06 09:45 | 803.77 | 7.344 | 53.4303 | 1
40 | 2023-03-06 10:00 | 803.76 | 7.3479 | 53.4595 | 1
41 | 2023-03-06 10:15 | 800.77 | 7.3137 | 53.4532 | 1
42 | 2023-03-06 10:30 | 801.76 | 7.2765 | 53.4174 | 1
43 | 2023-03-06 10:45 | 802.77 | 7.3025 | 53.3891 | 1
44 | 2023-03-06 11:00 | 803.76 | 7.3346 | 53.3972 | 1
45 | 2023-03-06 11:15 | 802.77 | 7.3143 | 53.3604 | 1
46 | 2023-03-06 11:30 | 803.76 | 7.3036 | 53.3004 | 1
47 | 2023-03-06 11:45 | 802.76 | 7.2972 | 53.3785 | 1
48 | 2023-03-06 12:00 | 803.76 | 7.3466 | 53.4077 | 1
49 | 2023-03-06 12:15 | 804.76 | 7.341 | 53.4329 | 1
50 | 2023-03-06 12:30 | 800.77 | 7.3638 | 53.4362 | 1
51 | 2023-03-06 12:45 | 801.77 | 7.3917 | 53.4263 | 1
52 | 2023-03-06 13:00 | 801.76 | 7.3883 | 53.4504 | 1
53 | 2023-03-06 13:15 | 802.77 | 7.4134 | 53.4626 | 1
54 | 2023-03-06 13:30 | 801.77 | 7.3552 | 53.3504 | 1
55 | 2023-03-06 13:45 | 803.76 | 7.3343 | 53.3265 | 1
56 | 2023-03-06 14:00 | 803.76 | 7.35 | 53.3951 | 1
57 | 2023-03-06 14:15 | 802.77 | 7.3472 | 53.3766 | 1
58 | 2023-03-06 14:30 | 803.76 | 7.3861 | 53.4097 | 1
59 | 2023-03-06 14:45 | 802.76 | 7.3591 | 53.3871 | 1
60 | 2023-03-06 15:00 | 802.77 | 7.3389 | 53.3542 | 1
61 | 2023-03-06 15:15 | 803.76 | 7.3255 | 53.3871 | 1
62 | 2023-03-06 15:30 | 802.76 | 7.2924 | 53.3849 | 1
63 | 2023-03-06 15:45 | 802.77 | 7.2881 | 53.3397 | 1
64 | 2023-03-06 16:00 | 803.76 | 7.306 | 53.3367 | 1
65 | 2023-03-06 16:15 | 873.6 | 7.1088 | 52.9275 | 1
66 | 2023-03-06 16:30 | 958.41 | 7.0895 | 52.8862 | 1
67 | 2023-03-06 16:45 | 942.44 | 7.1157 | 52.8933 | 1
68 | 2023-03-06 17:00 | 942.44 | 7.1119 | 52.8307 | 1
69 | 2023-03-06 17:15 | 942.45 | 7.048 | 52.6507 | 1
70 | 2023-03-06 17:30 | 945.43 | 6.9807 | 52.4836 | 1
71 | 2023-03-06 17:45 | 948.43 | 6.8995 | 52.2259 | 1
72 | 2023-03-06 18:00 | 955.42 | 7.062 | 52.4923 | 1
73 | 2023-03-06 18:15 | 942.44 | 7.2094 | 52.7987 | 1
74 | 2023-03-06 18:30 | 925.48 | 7.2995 | 53.0967 | 1
75 | 2023-03-06 18:45 | 921.5 | 7.2972 | 53.1748 | 1
76 | 2023-03-06 19:00 | 920.49 | 7.2898 | 53.1979 | 1
77 | 2023-03-06 19:15 | 926.49 | 7.2597 | 53.2033 | 1
78 | 2023-03-06 19:30 | 937.45 | 7.2703 | 53.2489 | 1
79 | 2023-03-06 19:45 | 935.46 | 7.2473 | 53.2904 | 1
80 | 2023-03-06 20:00 | 933.47 | 7.2256 | 53.258 | 1
81 | 2023-03-06 20:15 | 934.46 | 7.1888 | 53.1425 | 1
82 | 2023-03-06 20:30 | 936.46 | 7.1733 | 53.0979 | 1
83 | 2023-03-06 20:45 | 937.46 | 7.1677 | 53.1057 | 1
84 | 2023-03-06 21:00 | 936.45 | 7.111 | 53.039 | 1
85 | 2023-03-06 21:15 | 937.46 | 7.1035 | 52.9649 | 1
86 | 2023-03-06 21:30 | 938.45 | 7.0667 | 52.8839 | 1
87 | 2023-03-06 21:45 | 939.45 | 7.0757 | 52.8933 | 1
88 | 2023-03-06 22:00 | 936.46 | 7.0886 | 52.8856 | 1
89 | 2023-03-06 22:15 | 932.47 | 7.1394 | 53.0044 | 1
90 | 2023-03-06 22:30 | 917.5 | 7.1781 | 53.1157 | 1
91 | 2023-03-06 22:45 | 913.51 | 7.2067 | 53.18 | 1
92 | 2023-03-06 23:00 | 916.51 | 7.2146 | 53.2217 | 1
93 | 2023-03-06 23:15 | 910.52 | 7.203 | 53.2271 | 1
94 | 2023-03-06 23:30 | 913.51 | 7.1794 | 53.1884 | 1
95 | 2023-03-06 23:45 | 914.51 | 7.1619 | 53.1545 | 1
96 | 2023-03-07 | 914.51 | 7.1574 | 53.1367 | 1
97 | 2023-03-07 00:15 | 914.51 | 7.172 | 53.1282 | 1
98 | 2023-03-07 00:30 | 914.51 | 7.1682 | 53.1279 | 1
99 | 2023-03-07 00:45 | 915.51 | 7.1772 | 53.14 | 1
100 | 2023-03-07 01:00 | 915.51 | 7.1536 | 53.1264 | 1
101 | 2023-03-07 01:15 | 914.51 | 7.1525 | 53.1023 | 1
102 | 2023-03-07 01:30 | 915.51 | 7.1518 | 53.108 | 1
103 | 2023-03-07 01:45 | 914.51 | 7.1596 | 53.1109 | 1
104 | 2023-03-07 02:00 | 915.51 | 7.1841 | 53.124 | 1
105 | 2023-03-07 02:15 | 914.51 | 7.1561 | 53.1065 | 1
106 | 2023-03-07 02:30 | 915.51 | 7.1437 | 53.0963 | 1
107 | 2023-03-07 02:45 | 915.5 | 7.1523 | 53.1022 | 1
108 | 2023-03-07 03:00 | 915.51 | 7.145 | 53.0949 | 1
109 | 2023-03-07 03:15 | 915.51 | 7.1712 | 53.0884 | 1
110 | 2023-03-07 03:30 | 915.51 | 7.1587 | 53.0952 | 1
111 | 2023-03-07 03:45 | 915.51 | 7.1598 | 53.0856 | 1
112 | 2023-03-07 04:00 | 915.51 | 7.1609 | 53.0651 | 1
113 | 2023-03-07 04:15 | 915.5 | 7.1763 | 53.1362 | 1
114 | 2023-03-07 04:30 | 915.51 | 7.1751 | 53.1795 | 1
115 | 2023-03-07 04:45 | 913.51 | 7.1922 | 53.1778 | 1
116 | 2023-03-07 05:00 | 912.52 | 7.1701 | 53.1442 | 1
117 | 2023-03-07 05:15 | 914.51 | 7.1548 | 53.1656 | 1
118 | 2023-03-07 05:30 | 914.51 | 7.1733 | 53.2006 | 1
119 | 2023-03-07 05:45 | 913.51 | 7.1604 | 53.2028 | 1
120 | 2023-03-07 06:00 | 914.51 | 7.1905 | 53.1771 | 1
121 | 2023-03-07 06:15 | 914.51 | 7.1407 | 53.1336 | 1
122 | 2023-03-07 06:30 | 918.51 | 7.1746 | 53.1304 | 1
123 | 2023-03-07 06:45 | 913.51 | 7.1372 | 53.1512 | 1
124 | 2023-03-07 07:00 | 915.5 | 7.1497 | 53.1429 | 1
125 | 2023-03-07 07:15 | 916.51 | 7.1301 | 53.1207 | 1
126 | 2023-03-07 07:30 | 916.51 | 7.1351 | 53.1031 | 1
127 | 2023-03-07 07:45 | 916.5 | 7.1325 | 53.1122 | 1
128 | 2023-03-07 08:00 | 916.51 | 7.1142 | 53.0907 | 1
129 | 2023-03-07 08:15 | 915.51 | 7.0701 | 52.9922 | 1
130 | 2023-03-07 08:30 | 917.5 | 7.0551 | 52.9423 | 1
131 | 2023-03-07 08:45 | 922.49 | 7.0904 | 52.9468 | 1
132 | 2023-03-07 09:00 | 918.5 | 7.1002 | 52.9684 | 1
133 | 2023-03-07 09:15 | 923.49 | 7.1196 | 53.0016 | 1
134 | 2023-03-07 09:30 | 922.49 | 7.1198 | 53.0521 | 1
135 | 2023-03-07 09:45 | 920.5 | 7.1443 | 53.0871 | 1
136 | 2023-03-07 10:00 | 915.51 | 7.1531 | 53.0928 | 1
137 | 2023-03-07 10:15 | 916.5 | 7.2473 | 53.2778 | 1
138 | 2023-03-07 10:30 | 913.52 | 7.2473 | 53.3597 | 1
139 | 2023-03-07 10:45 | 910.51 | 7.203 | 53.293 | 1
140 | 2023-03-07 11:00 | 912.52 | 7.1514 | 53.1951 | 1
141 | 2023-03-07 11:15 | 915.51 | 7.2195 | 53.358 | 1
142 | 2023-03-07 11:30 | 912.51 | 7.2619 | 53.4441 | 1
143 | 2023-03-07 11:45 | 909.53 | 7.1772 | 53.3474 | 1
144 | 2023-03-07 12:00 | 912.51 | 7.151 | 53.2652 | 1
145 | 2023-03-07 12:15 | 913.51 | 7.1953 | 53.2765 | 1
146 | 2023-03-07 12:30 | 913.52 | 7.1714 | 53.3177 | 1
147 | 2023-03-07 12:45 | 912.51 | 7.1983 | 53.3405 | 1
148 | 2023-03-07 13:00 | 912.52 | 7.192 | 53.3372 | 1
149 | 2023-03-07 13:15 | 978.42 | 6.9996 | 52.9147 | 1
150 | 2023-03-07 13:30 | 1063.84 | 7.0054 | 52.8821 | 1
151 | 2023-03-07 13:45 | 1063.84 | 7.0054 | 52.8821 | 1
152 | 2023-03-07 14:00 | 1059.34 | 7.0198 | 52.8746 | 1
153 | 2023-03-07 14:15 | 1049.32 | 7.0431 | 52.9514 | 1
154 | 2023-03-07 14:30 | 1031.32 | 7.0792 | 52.9923 | 1
155 | 2023-03-07 14:45 | 1030.31 | 7.077 | 52.9919 | 1
156 | 2023-03-07 15:00 | 1030.31 | 7.0628 | 52.9684 | 1
157 | 2023-03-07 15:15 | 1030.31 | 7.0315 | 52.8957 | 1
158 | 2023-03-07 15:30 | 1032.32 | 7.0409 | 52.8584 | 1
159 | 2023-03-07 15:45 | 1031.32 | 7.0405 | 52.8616 | 1
160 | 2023-03-07 16:00 | 1033.31 | 7.0396 | 52.8497 | 1
161 | 2023-03-07 16:15 | 1041.32 | 7.0366 | 52.8473 | 1
162 | 2023-03-07 16:30 | 1044.32 | 7.0276 | 52.8533 | 1
163 | 2023-03-07 16:45 | 1047.33 | 7.0504 | 52.9106 | 1
164 | 2023-03-07 17:00 | 1046.32 | 7.0637 | 52.9436 | 1
165 | 2023-03-07 17:15 | 1045.33 | 7.0482 | 52.8883 | 1
166 | 2023-03-07 17:30 | 1045.32 | 7.016 | 52.8555 | 1
167 | 2023-03-07 17:45 | 1046.33 | 7.0108 | 52.8542 | 1
168 | 2023-03-07 18:00 | 1047.32 | 7.0039 | 52.8414 | 1
169 | 2023-03-07 18:15 | 1046.33 | 7.0052 | 52.8102 | 1
170 | 2023-03-07 18:30 | 1047.32 | 7.0117 | 52.7978 | 1
171 | 2023-03-07 18:45 | 1048.33 | 7.0134 | 52.8143 | 1
172 | 2023-03-07 19:00 | 1046.32 | 7.0014 | 52.8056 | 1
173 | 2023-03-07 19:15 | 1044.33 | 6.9932 | 52.7508 | 1
174 | 2023-03-07 19:30 | 1047.33 | 6.9932 | 52.7491 | 1
175 | 2023-03-07 19:45 | 1047.32 | 6.9807 | 52.6977 | 1
176 | 2023-03-07 20:00 | 1048.33 | 6.9672 | 52.6645 | 1
177 | 2023-03-07 20:15 | 1051.32 | 6.9435 | 52.6404 | 1
178 | 2023-03-07 20:30 | 1048.33 | 7.0226 | 52.7712 | 1
179 | 2023-03-07 20:45 | 1047.33 | 7.016 | 52.7862 | 1
180 | 2023-03-07 21:00 | 1048.32 | 6.9829 | 52.7282 | 1
181 | 2023-03-07 21:15 | 1045.33 | 6.9764 | 52.7459 | 1
182 | 2023-03-07 21:30 | 1047.32 | 7.0246 | 52.7969 | 1
183 | 2023-03-07 21:45 | 1047.32 | 7.0155 | 52.8474 | 1
184 | 2023-03-07 22:00 | 1041.33 | 7.0428 | 52.8985 | 1
185 | 2023-03-07 22:15 | 1045.32 | 7.0413 | 52.9073 | 1
186 | 2023-03-07 22:30 | 1048.28 | 7.0042 | 52.8675 | 1
187 | 2023-03-07 22:45 | 1049.62 | 7.0156 | 52.9119 | 1
188 | 2023-03-07 23:00 | 1048.17 | 7.0448 | 53.0211 | 1
189 | 2023-03-07 23:15 | 1044.9 | 7.0374 | 53.035 | 1
190 | 2023-03-07 23:30 | 1043.51 | 7.0193 | 53.0233 | 1
191 | 2023-03-07 23:45 | 1044.14 | 7.0069 | 52.9999 | 1
192 | 2023-03-08 | 1043.72 | 7.0003 | 53.0068 | 1

## Composition
DateTime | Unit | H2 | N2 | CO2 | C1 | C2 | C3 | iC4 | nC4 | iC5 | nC5 | neoC5 | C6 | He
2023-03-06 00:15 | cmol_mol-1 | 0 | 10.5476 | 1.0243 | 83.4961 | 3.8399 | 0.6835 | 0.1233 | 0.1144 | 0.0321 | 0.0251 | 0.0024 | 0.0614 | 0.05
2023-03-06 00:30 | cmol_mol-1 | 0 | 10.5451 | 1.0294 | 83.4841 | 3.8464 | 0.6856 | 0.1241 | 0.1148 | 0.0318 | 0.025 | 0.0024 | 0.0614 | 0.05
2023-03-06 00:45 | cmol_mol-1 | 0 | 10.5525 | 1.0317 | 83.4645 | 3.8539 | 0.6869 | 0.1248 | 0.1153 | 0.0319 | 0.0248 | 0.0025 | 0.0614 | 0.05
2023-03-06 01:00 | cmol_mol-1 | 0 | 10.5332 | 1.0411 | 83.4422 | 3.8762 | 0.6942 | 0.1261 | 0.1163 | 0.032 | 0.0249 | 0.0023 | 0.0616 | 0.05
2023-03-06 01:15 | cmol_mol-1 | 0 | 10.5302 | 1.0378 | 83.4704 | 3.8585 | 0.6909 | 0.1253 | 0.1159 | 0.0321 | 0.0248 | 0.0025 | 0.0616 | 0.05
2023-03-06 01:30 | cmol_mol-1 | 0 | 10.5647 | 1.0383 | 83.4388 | 3.8547 | 0.6913 | 0.1252 | 0.1161 | 0.0321 | 0.0251 | 0.0025 | 0.0613 | 0.05
2023-03-06 01:45 | cmol_mol-1 | 0 | 10.5635 | 1.0359 | 83.4642 | 3.8376 | 0.6873 | 0.1246 | 0.1161 | 0.0321 | 0.025 | 0.0025 | 0.0611 | 0.05
2023-03-06 02:00 | cmol_mol-1 | 0 | 10.5533 | 1.0337 | 83.4818 | 3.8322 | 0.6861 | 0.1248 | 0.1171 | 0.032 | 0.0251 | 0.0026 | 0.0612 | 0.05
2023-03-06 02:15 | cmol_mol-1 | 0 | 10.5642 | 1.024 | 83.5151 | 3.8065 | 0.6791 | 0.1233 | 0.1166 | 0.0323 | 0.0252 | 0.0027 | 0.0608 | 0.05
2023-03-06 02:30 | cmol_mol-1 | 0 | 10.5539 | 1.0156 | 83.5615 | 3.7859 | 0.6737 | 0.122 | 0.1159 | 0.0324 | 0.0254 | 0.0026 | 0.061 | 0.05
2023-03-06 02:45 | cmol_mol-1 | 0 | 10.5664 | 1.0025 | 83.6005 | 3.7588 | 0.6654 | 0.1201 | 0.115 | 0.0324 | 0.0254 | 0.0025 | 0.0608 | 0.05
2023-03-06 03:00 | cmol_mol-1 | 0 | 10.5376 | 0.9892 | 83.6811 | 3.732 | 0.6576 | 0.1179 | 0.1141 | 0.032 | 0.0255 | 0.0025 | 0.0606 | 0.05
2023-03-06 03:15 | cmol_mol-1 | 0 | 10.5361 | 0.9743 | 83.7414 | 3.7021 | 0.6481 | 0.1153 | 0.1127 | 0.0314 | 0.0253 | 0.0026 | 0.0608 | 0.05
2023-03-06 03:30 | cmol_mol-1 | 0 | 10.5322 | 0.965 | 83.7869 | 3.6789 | 0.6417 | 0.1137 | 0.1122 | 0.0308 | 0.025 | 0.0026 | 0.0608 | 0.05
2023-03-06 03:45 | cmol_mol-1 | 0 | 10.5285 | 0.9577 | 83.8203 | 3.6635 | 0.6372 | 0.1119 | 0.1114 | 0.0308 | 0.0252 | 0.0026 | 0.0608 | 0.05
2023-03-06 04:00 | cmol_mol-1 | 0 | 10.5072 | 0.9601 | 83.8489 | 3.6562 | 0.636 | 0.1114 | 0.1114 | 0.0306 | 0.0248 | 0.0025 | 0.0609 | 0.05
2023-03-06 04:15 | cmol_mol-1 | 0 | 10.5317 | 0.9657 | 83.795 | 3.6734 | 0.6409 | 0.1118 | 0.1124 | 0.0307 | 0.025 | 0.0025 | 0.061 | 0.05
2023-03-06 04:30 | cmol_mol-1 | 0 | 10.5268 | 0.9683 | 83.7899 | 3.6782 | 0.6423 | 0.1123 | 0.113 | 0.0307 | 0.0249 | 0.0025 | 0.061 | 0.05
2023-03-06 04:45 | cmol_mol-1 | 0 | 10.5469 | 0.9717 | 83.7406 | 3.6976 | 0.6465 | 0.1133 | 0.1139 | 0.0311 | 0.0249 | 0.0024 | 0.0611 | 0.05
2023-03-06 05:00 | cmol_mol-1 | 0 | 10.5449 | 0.9748 | 83.7285 | 3.7054 | 0.6482 | 0.1135 | 0.1146 | 0.0309 | 0.0252 | 0.0025 | 0.0615 | 0.05
2023-03-06 05:15 | cmol_mol-1 | 0 | 10.5589 | 0.9797 | 83.7038 | 3.7086 | 0.6495 | 0.1138 | 0.1152 | 0.0314 | 0.0254 | 0.0025 | 0.0612 | 0.05
2023-03-06 05:30 | cmol_mol-1 | 0 | 10.5518 | 0.9841 | 83.6971 | 3.715 | 0.6519 | 0.1138 | 0.1152 | 0.0316 | 0.0253 | 0.0025 | 0.0618 | 0.05
2023-03-06 05:45 | cmol_mol-1 | 0 | 10.5518 | 0.9893 | 83.6779 | 3.7253 | 0.6546 | 0.1142 | 0.1152 | 0.0315 | 0.0255 | 0.0026 | 0.0622 | 0.05
2023-03-06 06:00 | cmol_mol-1 | 0 | 10.5599 | 0.9941 | 83.6544 | 3.7321 | 0.6569 | 0.1149 | 0.1156 | 0.0317 | 0.0257 | 0.0025 | 0.0621 | 0.05
2023-03-06 06:15 | cmol_mol-1 | 0 | 10.5493 | 0.9995 | 83.6409 | 3.7459 | 0.6601 | 0.1156 | 0.1165 | 0.0319 | 0.0259 | 0.0025 | 0.062 | 0.05
2023-03-06 06:30 | cmol_mol-1 | 0 | 10.5497 | 1.0037 | 83.6175 | 3.761 | 0.6638 | 0.1159 | 0.1168 | 0.0319 | 0.026 | 0.0025 | 0.0612 | 0.05
2023-03-06 06:45 | cmol_mol-1 | 0 | 10.5529 | 1.0226 | 83.5542 | 3.7866 | 0.6747 | 0.118 | 0.1188 | 0.0325 | 0.0263 | 0.0025 | 0.0611 | 0.05
2023-03-06 07:00 | cmol_mol-1 | 0 | 10.5418 | 1.0393 | 83.5191 | 3.7996 | 0.6828 | 0.1207 | 0.1221 | 0.0331 | 0.0268 | 0.0026 | 0.0621 | 0.05
2023-03-06 07:15 | cmol_mol-1 | 0 | 10.5312 | 1.0471 | 83.5056 | 3.8089 | 0.6856 | 0.1217 | 0.1228 | 0.0341 | 0.0276 | 0.0026 | 0.0628 | 0.05
2023-03-06 07:30 | cmol_mol-1 | 0 | 10.5244 | 1.0619 | 83.4304 | 3.8588 | 0.6985 | 0.124 | 0.1227 | 0.0347 | 0.028 | 0.0026 | 0.064 | 0.05
2023-03-06 07:45 | cmol_mol-1 | 0 | 10.5337 | 1.0789 | 83.3354 | 3.9077 | 0.713 | 0.1276 | 0.1243 | 0.0343 | 0.0277 | 0.0025 | 0.065 | 0.05
2023-03-06 08:00 | cmol_mol-1 | 0 | 10.5359 | 1.0919 | 83.2753 | 3.9406 | 0.7207 | 0.1299 | 0.1252 | 0.0347 | 0.0277 | 0.0026 | 0.0654 | 0.05
2023-03-06 08:15 | cmol_mol-1 | 0 | 10.5314 | 1.1222 | 83.1868 | 3.9844 | 0.7332 | 0.132 | 0.1265 | 0.0354 | 0.0286 | 0.0026 | 0.0669 | 0.05
2023-03-06 08:30 | cmol_mol-1 | 0 | 10.5239 | 1.1617 | 83.071 | 4.0418 | 0.7507 | 0.1356 | 0.1283 | 0.0359 | 0.0288 | 0.0027 | 0.0697 | 0.05
2023-03-06 08:45 | cmol_mol-1 | 0 | 10.4992 | 1.1875 | 83.0676 | 4.035 | 0.7555 | 0.1368 | 0.1302 | 0.0356 | 0.0288 | 0.0029 | 0.0708 | 0.05
2023-03-06 09:00 | cmol_mol-1 | 0 | 10.4857 | 1.1662 | 83.2074 | 3.9563 | 0.7354 | 0.1323 | 0.1293 | 0.0352 | 0.0287 | 0.0027 | 0.0708 | 0.05
2023-03-06 09:15 | cmol_mol-1 | 0 | 10.4723 | 1.1253 | 83.3819 | 3.8749 | 0.71 | 0.1258 | 0.1266 | 0.0341 | 0.028 | 0.0028 | 0.0684 | 0.05
2023-03-06 09:30 | cmol_mol-1 | 0 | 10.5132 | 1.1094 | 83.4147 | 3.8324 | 0.6992 | 0.1227 | 0.1263 | 0.034 | 0.0282 | 0.0029 | 0.0672 | 0.05
2023-03-06 09:45 | cmol_mol-1 | 0 | 10.5028 | 1.1033 | 83.4638 | 3.8063 | 0.694 | 0.1209 | 0.1269 | 0.0341 | 0.0283 | 0.0027 | 0.0668 | 0.05
2023-03-06 10:00 | cmol_mol-1 | 0 | 10.4944 | 1.0935 | 83.5062 | 3.7891 | 0.6884 | 0.1201 | 0.1267 | 0.034 | 0.0282 | 0.0028 | 0.0665 | 0.05
2023-03-06 10:15 | cmol_mol-1 | 0 | 10.4951 | 1.0931 | 83.5052 | 3.7906 | 0.6875 | 0.12 | 0.1269 | 0.0341 | 0.0284 | 0.0029 | 0.0663 | 0.05
2023-03-06 10:30 | cmol_mol-1 | 0 | 10.5064 | 1.0937 | 83.4911 | 3.793 | 0.6876 | 0.12 | 0.1269 | 0.0344 | 0.0284 | 0.0028 | 0.0658 | 0.05
2023-03-06 10:45 | cmol_mol-1 | 0 | 10.4903 | 1.096 | 83.499 | 3.7975 | 0.689 | 0.1199 | 0.1265 | 0.0346 | 0.0287 | 0.0028 | 0.0656 | 0.05
2023-03-06 11:00 | cmol_mol-1 | 0 | 10.4855 | 1.095 | 83.5086 | 3.7948 | 0.6887 | 0.1198 | 0.126 | 0.0345 | 0.0287 | 0.0027 | 0.0657 | 0.05
2023-03-06 11:15 | cmol_mol-1 | 0 | 10.4814 | 1.0934 | 83.5191 | 3.7916 | 0.6881 | 0.1201 | 0.1265 | 0.0344 | 0.0287 | 0.0028 | 0.064 | 0.05
2023-03-06 11:30 | cmol_mol-1 | 0 | 10.4747 | 1.0996 | 83.5144 | 3.7958 | 0.6888 | 0.1201 | 0.1268 | 0.0347 | 0.029 | 0.0029 | 0.0632 | 0.05
2023-03-06 11:45 | cmol_mol-1 | 0 | 10.4664 | 1.1035 | 83.5115 | 3.8024 | 0.6907 | 0.1202 | 0.1258 | 0.0345 | 0.0286 | 0.0028 | 0.0635 | 0.05
2023-03-06 12:00 | cmol_mol-1 | 0 | 10.4685 | 1.1032 | 83.5109 | 3.7983 | 0.692 | 0.1217 | 0.1275 | 0.0344 | 0.0287 | 0.0029 | 0.062 | 0.05
2023-03-06 12:15 | cmol_mol-1 | 0 | 10.497 | 1.1073 | 83.4764 | 3.7994 | 0.6893 | 0.1219 | 0.1295 | 0.0351 | 0.0292 | 0.0028 | 0.0621 | 0.05
2023-03-06 12:30 | cmol_mol-1 | 0 | 10.4453 | 1.1133 | 83.5104 | 3.8102 | 0.6901 | 0.1206 | 0.1279 | 0.0363 | 0.0302 | 0.0028 | 0.0629 | 0.05
2023-03-06 12:45 | cmol_mol-1 | 0 | 10.4076 | 1.1019 | 83.5777 | 3.8008 | 0.687 | 0.1196 | 0.125 | 0.035 | 0.0297 | 0.0029 | 0.0629 | 0.05
2023-03-06 13:00 | cmol_mol-1 | 0 | 10.45 | 1.0792 | 83.5962 | 3.7767 | 0.6769 | 0.1191 | 0.1239 | 0.0342 | 0.0283 | 0.0027 | 0.0626 | 0.05
2023-03-06 13:15 | cmol_mol-1 | 0 | 10.4373 | 1.0656 | 83.6632 | 3.7534 | 0.666 | 0.1167 | 0.1215 | 0.0335 | 0.0278 | 0.0026 | 0.0624 | 0.05
2023-03-06 13:30 | cmol_mol-1 | 0 | 10.4409 | 1.0599 | 83.7024 | 3.73 | 0.6576 | 0.1145 | 0.1192 | 0.0331 | 0.0274 | 0.0027 | 0.0623 | 0.05
2023-03-06 13:45 | cmol_mol-1 | 0 | 10.443 | 1.0556 | 83.7232 | 3.7191 | 0.6527 | 0.1132 | 0.1181 | 0.0327 | 0.0276 | 0.0027 | 0.0621 | 0.05
2023-03-06 14:00 | cmol_mol-1 | 0 | 10.4425 | 1.0534 | 83.7385 | 3.7138 | 0.6504 | 0.1119 | 0.1162 | 0.0322 | 0.0267 | 0.0026 | 0.0617 | 0.05
2023-03-06 14:15 | cmol_mol-1 | 0 | 10.431 | 1.0551 | 83.7457 | 3.7139 | 0.653 | 0.1123 | 0.1163 | 0.0316 | 0.0265 | 0.0024 | 0.0621 | 0.05
2023-03-06 14:30 | cmol_mol-1 | 0 | 10.4405 | 1.0434 | 83.7634 | 3.7029 | 0.6499 | 0.1121 | 0.1162 | 0.0312 | 0.0262 | 0.0025 | 0.0617 | 0.05
2023-03-06 14:45 | cmol_mol-1 | 0 | 10.4426 | 1.0357 | 83.7859 | 3.6933 | 0.645 | 0.1114 | 0.1159 | 0.0308 | 0.0255 | 0.0027 | 0.0614 | 0.05
2023-03-06 15:00 | cmol_mol-1 | 0 | 10.4512 | 1.0169 | 83.8304 | 3.6765 | 0.6348 | 0.1095 | 0.1134 | 0.0298 | 0.0248 | 0.0027 | 0.0601 | 0.05
2023-03-06 15:15 | cmol_mol-1 | 0 | 10.4427 | 1.0196 | 83.8258 | 3.6841 | 0.6368 | 0.1095 | 0.1134 | 0.0302 | 0.025 | 0.0028 | 0.0603 | 0.05
2023-03-06 15:30 | cmol_mol-1 | 0 | 10.4348 | 1.0116 | 83.8482 | 3.6785 | 0.6365 | 0.1094 | 0.1134 | 0.0302 | 0.0248 | 0.0026 | 0.06 | 0.05
2023-03-06 15:45 | cmol_mol-1 | 0 | 10.4526 | 1.0005 | 83.8568 | 3.6684 | 0.633 | 0.1089 | 0.1132 | 0.0299 | 0.0247 | 0.0026 | 0.0594 | 0.05
2023-03-06 16:00 | cmol_mol-1 | 0 | 10.4768 | 0.9918 | 83.862 | 3.6547 | 0.6275 | 0.108 | 0.1131 | 0.03 | 0.0248 | 0.0026 | 0.0588 | 0.05
2023-03-06 16:15 | cmol_mol-1 | 0 | 10.4722 | 0.999 | 83.8508 | 3.6576 | 0.6309 | 0.1082 | 0.1144 | 0.0302 | 0.025 | 0.0027 | 0.0589 | 0.05
2023-03-06 16:30 | cmol_mol-1 | 0 | 10.4831 | 0.9958 | 83.8581 | 3.6449 | 0.6297 | 0.1082 | 0.1142 | 0.0304 | 0.0252 | 0.0028 | 0.0577 | 0.05
2023-03-06 16:45 | cmol_mol-1 | 0 | 10.4823 | 0.9984 | 83.8455 | 3.6534 | 0.632 | 0.1087 | 0.1139 | 0.0308 | 0.0257 | 0.0026 | 0.0567 | 0.05
2023-03-06 17:00 | cmol_mol-1 | 0 | 10.5028 | 1.0158 | 83.7705 | 3.6705 | 0.6439 | 0.1115 | 0.1169 | 0.0315 | 0.0264 | 0.0026 | 0.0576 | 0.05
2023-03-06 17:15 | cmol_mol-1 | 0 | 10.4697 | 1.0389 | 83.7606 | 3.677 | 0.648 | 0.1132 | 0.1207 | 0.0324 | 0.0273 | 0.0028 | 0.0595 | 0.05
2023-03-06 17:30 | cmol_mol-1 | 0 | 10.4691 | 1.0678 | 83.6501 | 3.7336 | 0.6644 | 0.116 | 0.1222 | 0.034 | 0.0287 | 0.0028 | 0.0615 | 0.05
2023-03-06 17:45 | cmol_mol-1 | 0 | 10.4713 | 1.0643 | 83.6509 | 3.7339 | 0.6659 | 0.1156 | 0.1215 | 0.0336 | 0.0288 | 0.0026 | 0.0617 | 0.05
2023-03-06 18:00 | cmol_mol-1 | 0 | 10.4852 | 1.066 | 83.627 | 3.7371 | 0.6695 | 0.1165 | 0.1222 | 0.0332 | 0.0283 | 0.0026 | 0.0625 | 0.05
2023-03-06 18:15 | cmol_mol-1 | 0 | 10.4879 | 1.0602 | 83.631 | 3.7366 | 0.6673 | 0.1166 | 0.1228 | 0.0333 | 0.028 | 0.0027 | 0.0635 | 0.05
2023-03-06 18:30 | cmol_mol-1 | 0 | 10.9741 | 1.0408 | 83.208 | 3.7093 | 0.6589 | 0.1141 | 0.119 | 0.0328 | 0.0277 | 0.0026 | 0.0627 | 0.05
2023-03-06 18:45 | cmol_mol-1 | 0 | 11.251 | 1.0341 | 82.9443 | 3.7044 | 0.6596 | 0.1147 | 0.1185 | 0.032 | 0.0267 | 0.0025 | 0.0623 | 0.05
2023-03-06 19:00 | cmol_mol-1 | 0 | 11.2597 | 1.0643 | 82.8092 | 3.77 | 0.6801 | 0.1194 | 0.1219 | 0.0324 | 0.0269 | 0.0026 | 0.0636 | 0.05
2023-03-06 19:15 | cmol_mol-1 | 0 | 11.326 | 1.0376 | 82.7343 | 3.8004 | 0.6848 | 0.1193 | 0.1227 | 0.0322 | 0.0266 | 0.0025 | 0.0636 | 0.05
2023-03-06 19:30 | cmol_mol-1 | 0 | 11.2971 | 1.0436 | 82.78 | 3.783 | 0.6812 | 0.1185 | 0.1222 | 0.0326 | 0.0267 | 0.0025 | 0.0626 | 0.05
2023-03-06 19:45 | cmol_mol-1 | 0 | 11.3395 | 1.0961 | 82.5536 | 3.8664 | 0.7121 | 0.1257 | 0.1276 | 0.0341 | 0.0278 | 0.0026 | 0.0645 | 0.05
2023-03-06 20:00 | cmol_mol-1 | 0 | 11.2849 | 1.1186 | 82.545 | 3.8916 | 0.7204 | 0.1276 | 0.1298 | 0.0346 | 0.0284 | 0.0027 | 0.0664 | 0.05
2023-03-06 20:15 | cmol_mol-1 | 0 | 11.2653 | 1.1313 | 82.5362 | 3.8976 | 0.7241 | 0.1282 | 0.1317 | 0.0357 | 0.0295 | 0.0027 | 0.0677 | 0.05
2023-03-06 20:30 | cmol_mol-1 | 0 | 11.3035 | 1.1309 | 82.4677 | 3.9249 | 0.7282 | 0.1285 | 0.1313 | 0.0353 | 0.0295 | 0.0026 | 0.0675 | 0.05
2023-03-06 20:45 | cmol_mol-1 | 0 | 11.3206 | 1.1331 | 82.3735 | 3.9903 | 0.7391 | 0.1302 | 0.1294 | 0.0352 | 0.0288 | 0.0025 | 0.0674 | 0.05
2023-03-06 21:00 | cmol_mol-1 | 0 | 11.3301 | 1.1256 | 82.3722 | 3.997 | 0.7363 | 0.13 | 0.1276 | 0.0347 | 0.028 | 0.0023 | 0.0663 | 0.05
2023-03-06 21:15 | cmol_mol-1 | 0 | 11.3294 | 1.1221 | 82.3868 | 3.9998 | 0.7305 | 0.1281 | 0.1246 | 0.0338 | 0.0274 | 0.0022 | 0.0651 | 0.05
2023-03-06 21:30 | cmol_mol-1 | 0 | 10.6059 | 1.1133 | 83.0709 | 4.0543 | 0.7311 | 0.1298 | 0.1208 | 0.0329 | 0.0259 | 0.002 | 0.0631 | 0.05
2023-03-06 21:45 | cmol_mol-1 | 0 | 10.5869 | 1.1579 | 82.8959 | 4.1696 | 0.759 | 0.1346 | 0.1231 | 0.0331 | 0.0256 | 0.0022 | 0.0622 | 0.05
2023-03-06 22:00 | cmol_mol-1 | 0 | 10.5193 | 1.1701 | 82.9988 | 4.1203 | 0.7559 | 0.1345 | 0.126 | 0.0339 | 0.0266 | 0.0026 | 0.0622 | 0.05
2023-03-06 22:15 | cmol_mol-1 | 0 | 10.5425 | 1.137 | 83.048 | 4.085 | 0.7506 | 0.1342 | 0.1257 | 0.0339 | 0.0262 | 0.0028 | 0.0643 | 0.05
2023-03-06 22:30 | cmol_mol-1 | 0 | 10.4605 | 1.1321 | 83.1585 | 4.054 | 0.7499 | 0.1338 | 0.1286 | 0.0348 | 0.0277 | 0.0029 | 0.0672 | 0.05
2023-03-06 22:45 | cmol_mol-1 | 0 | 10.5475 | 1.13 | 83.1544 | 3.9803 | 0.741 | 0.1306 | 0.1321 | 0.0355 | 0.0284 | 0.0034 | 0.067 | 0.05
2023-03-06 23:00 | cmol_mol-1 | 0 | 10.5179 | 1.1534 | 83.2096 | 3.9565 | 0.7267 | 0.1285 | 0.1269 | 0.0341 | 0.027 | 0.0035 | 0.0661 | 0.05
2023-03-06 23:15 | cmol_mol-1 | 0 | 10.5845 | 1.1399 | 83.2815 | 3.8798 | 0.696 | 0.1242 | 0.1194 | 0.0322 | 0.0251 | 0.0038 | 0.0634 | 0.05
2023-03-06 23:30 | cmol_mol-1 | 0 | 10.4013 | 1.1433 | 83.445 | 3.8888 | 0.6997 | 0.1275 | 0.1191 | 0.0324 | 0.0252 | 0.0039 | 0.0639 | 0.05
2023-03-06 23:45 | cmol_mol-1 | 0 | 10.4202 | 1.1428 | 83.4222 | 3.8879 | 0.7015 | 0.1282 | 0.1195 | 0.0328 | 0.0254 | 0.004 | 0.0654 | 0.05
2023-03-07 | cmol_mol-1 | 0 | 10.3789 | 1.1589 | 83.463 | 3.8852 | 0.6951 | 0.1264 | 0.1172 | 0.0319 | 0.0247 | 0.0039 | 0.0649 | 0.05
2023-03-07 00:15 | cmol_mol-1 | 0 | 10.3749 | 1.1525 | 83.4731 | 3.8837 | 0.6956 | 0.1277 | 0.1168 | 0.0321 | 0.0248 | 0.004 | 0.0649 | 0.05
2023-03-07 00:30 | cmol_mol-1 | 0 | 10.4012 | 1.1459 | 83.4859 | 3.8608 | 0.6887 | 0.1267 | 0.1155 | 0.0316 | 0.0246 | 0.0039 | 0.0653 | 0.05
2023-03-07 00:45 | cmol_mol-1 | 0 | 10.3629 | 1.1496 | 83.6225 | 3.7914 | 0.6656 | 0.122 | 0.1117 | 0.0306 | 0.0239 | 0.0045 | 0.0653 | 0.05
2023-03-07 01:00 | cmol_mol-1 | 0 | 10.4282 | 1.1713 | 83.4159 | 3.8673 | 0.694 | 0.1267 | 0.1177 | 0.0322 | 0.0251 | 0.0041 | 0.0675 | 0.05
2023-03-07 01:15 | cmol_mol-1 | 0 | 10.404 | 1.1706 | 83.3891 | 3.9096 | 0.7023 | 0.1277 | 0.1184 | 0.0323 | 0.0251 | 0.0039 | 0.0671 | 0.05
2023-03-07 01:30 | cmol_mol-1 | 0 | 10.3885 | 1.1902 | 83.3471 | 3.9393 | 0.7086 | 0.1288 | 0.1191 | 0.0322 | 0.0254 | 0.0039 | 0.0669 | 0.05
2023-03-07 01:45 | cmol_mol-1 | 0 | 10.4084 | 1.1919 | 83.2738 | 3.9749 | 0.7204 | 0.1307 | 0.1212 | 0.0328 | 0.0254 | 0.0036 | 0.0669 | 0.05
2023-03-07 02:00 | cmol_mol-1 | 0 | 10.4686 | 1.1688 | 83.3052 | 3.9246 | 0.7075 | 0.1273 | 0.1199 | 0.0325 | 0.0255 | 0.0036 | 0.0665 | 0.05
2023-03-07 02:15 | cmol_mol-1 | 0 | 10.4496 | 1.1545 | 83.3654 | 3.9058 | 0.7019 | 0.1264 | 0.1194 | 0.0324 | 0.0255 | 0.0036 | 0.0655 | 0.05
2023-03-07 02:30 | cmol_mol-1 | 0 | 10.4707 | 1.1269 | 83.3971 | 3.891 | 0.6959 | 0.1242 | 0.1188 | 0.0319 | 0.0254 | 0.0034 | 0.0645 | 0.05
2023-03-07 02:45 | cmol_mol-1 | 0 | 10.423 | 1.1111 | 83.4814 | 3.8876 | 0.6866 | 0.1219 | 0.1151 | 0.0311 | 0.0242 | 0.0036 | 0.0644 | 0.05
2023-03-07 03:00 | cmol_mol-1 | 0 | 10.4218 | 1.1119 | 83.4463 | 3.9233 | 0.6903 | 0.1229 | 0.1133 | 0.0306 | 0.0233 | 0.0034 | 0.0628 | 0.05
2023-03-07 03:15 | cmol_mol-1 | 0 | 10.4172 | 1.1073 | 83.4559 | 3.9224 | 0.6902 | 0.1248 | 0.1126 | 0.0305 | 0.0232 | 0.0034 | 0.0624 | 0.05
2023-03-07 03:30 | cmol_mol-1 | 0 | 10.3949 | 1.1225 | 83.479 | 3.9015 | 0.6916 | 0.1244 | 0.1154 | 0.0311 | 0.0239 | 0.0035 | 0.0622 | 0.05
2023-03-07 03:45 | cmol_mol-1 | 0 | 10.4354 | 1.1186 | 83.3452 | 3.9733 | 0.7099 | 0.1273 | 0.1177 | 0.0319 | 0.0244 | 0.0032 | 0.0632 | 0.05
2023-03-07 04:00 | cmol_mol-1 | 0 | 10.5549 | 1.1162 | 83.0423 | 4.1166 | 0.7444 | 0.133 | 0.1206 | 0.0324 | 0.0245 | 0.0028 | 0.0624 | 0.05
2023-03-07 04:15 | cmol_mol-1 | 0 | 10.5028 | 1.1182 | 83.1954 | 4.0312 | 0.7284 | 0.131 | 0.1202 | 0.0326 | 0.0249 | 0.003 | 0.0624 | 0.05
2023-03-07 04:30 | cmol_mol-1 | 0 | 10.4442 | 1.1363 | 83.2631 | 4.0026 | 0.7257 | 0.1306 | 0.1214 | 0.0331 | 0.0258 | 0.0032 | 0.064 | 0.05
2023-03-07 04:45 | cmol_mol-1 | 0 | 10.4442 | 1.1491 | 83.2412 | 4.0118 | 0.7258 | 0.1309 | 0.1207 | 0.0329 | 0.0255 | 0.0031 | 0.0647 | 0.05
2023-03-07 05:00 | cmol_mol-1 | 0 | 10.4467 | 1.1578 | 83.2317 | 4.0127 | 0.7234 | 0.1307 | 0.1198 | 0.0326 | 0.0254 | 0.0031 | 0.0661 | 0.05
2023-03-07 05:15 | cmol_mol-1 | 0 | 10.4038 | 1.1617 | 83.2521 | 4.0271 | 0.7271 | 0.1315 | 0.1203 | 0.0326 | 0.0254 | 0.003 | 0.0654 | 0.05
2023-03-07 05:30 | cmol_mol-1 | 0 | 10.4205 | 1.1708 | 83.2192 | 4.0326 | 0.7289 | 0.1309 | 0.1208 | 0.0327 | 0.0255 | 0.003 | 0.065 | 0.05
2023-03-07 05:45 | cmol_mol-1 | 0 | 10.4303 | 1.1679 | 83.2157 | 4.0292 | 0.7281 | 0.1308 | 0.1211 | 0.0329 | 0.0259 | 0.0031 | 0.0651 | 0.05
2023-03-07 06:00 | cmol_mol-1 | 0 | 10.4353 | 1.1748 | 83.174 | 4.0469 | 0.7354 | 0.1322 | 0.1227 | 0.0336 | 0.0263 | 0.003 | 0.0658 | 0.05
2023-03-07 06:15 | cmol_mol-1 | 0 | 10.4929 | 1.1624 | 83.1351 | 4.0412 | 0.7354 | 0.1324 | 0.1227 | 0.0335 | 0.0261 | 0.003 | 0.0654 | 0.05
2023-03-07 06:30 | cmol_mol-1 | 0 | 10.4824 | 1.1579 | 83.1496 | 4.0414 | 0.7366 | 0.1327 | 0.1229 | 0.0333 | 0.026 | 0.0029 | 0.0641 | 0.05
2023-03-07 06:45 | cmol_mol-1 | 0 | 10.4669 | 1.1686 | 83.1389 | 4.0507 | 0.7407 | 0.1333 | 0.1239 | 0.0339 | 0.0265 | 0.0029 | 0.0636 | 0.05
2023-03-07 07:00 | cmol_mol-1 | 0 | 10.4827 | 1.1541 | 83.1368 | 4.0539 | 0.7391 | 0.1337 | 0.123 | 0.0335 | 0.0264 | 0.0028 | 0.064 | 0.05
2023-03-07 07:15 | cmol_mol-1 | 0 | 10.5237 | 1.1446 | 83.0462 | 4.0983 | 0.7495 | 0.1357 | 0.124 | 0.0343 | 0.0266 | 0.0027 | 0.0645 | 0.05
2023-03-07 07:30 | cmol_mol-1 | 0 | 10.4947 | 1.1681 | 82.9664 | 4.1736 | 0.7641 | 0.1355 | 0.1232 | 0.0335 | 0.0257 | 0.0026 | 0.0626 | 0.05
2023-03-07 07:45 | cmol_mol-1 | 0 | 10.5001 | 1.1651 | 83.0093 | 4.1301 | 0.7599 | 0.135 | 0.1247 | 0.0337 | 0.0262 | 0.0028 | 0.063 | 0.05
2023-03-07 08:00 | cmol_mol-1 | 0 | 10.5388 | 1.1568 | 82.9758 | 4.1295 | 0.762 | 0.1358 | 0.1255 | 0.0342 | 0.0262 | 0.0028 | 0.0627 | 0.05
2023-03-07 08:15 | cmol_mol-1 | 0 | 10.5361 | 1.1436 | 83.0154 | 4.1133 | 0.7569 | 0.135 | 0.1253 | 0.0339 | 0.0259 | 0.0026 | 0.0619 | 0.05
2023-03-07 08:30 | cmol_mol-1 | 0 | 10.5467 | 1.1031 | 83.1415 | 4.0425 | 0.7367 | 0.1327 | 0.1243 | 0.0334 | 0.0258 | 0.0027 | 0.0606 | 0.05
2023-03-07 08:45 | cmol_mol-1 | 0 | 10.552 | 1.0612 | 83.3098 | 3.9461 | 0.7099 | 0.1284 | 0.1222 | 0.0332 | 0.0256 | 0.0027 | 0.0588 | 0.05
2023-03-07 09:00 | cmol_mol-1 | 0 | 10.5613 | 1.0132 | 83.486 | 3.846 | 0.6823 | 0.1239 | 0.1179 | 0.0329 | 0.0252 | 0.0027 | 0.0585 | 0.05
2023-03-07 09:15 | cmol_mol-1 | 0 | 10.6004 | 0.984 | 83.5636 | 3.7862 | 0.6646 | 0.1202 | 0.1134 | 0.0328 | 0.025 | 0.0025 | 0.0572 | 0.05
2023-03-07 09:30 | cmol_mol-1 | 0 | 10.5724 | 0.9775 | 83.6128 | 3.776 | 0.663 | 0.1197 | 0.1116 | 0.0325 | 0.0253 | 0.0025 | 0.0569 | 0.05
2023-03-07 09:45 | cmol_mol-1 | 0 | 10.5486 | 1.1274 | 83.0363 | 4.1038 | 0.7536 | 0.1338 | 0.1237 | 0.0345 | 0.0269 | 0.0025 | 0.0588 | 0.05
2023-03-07 10:00 | cmol_mol-1 | 0 | 10.5275 | 1.2008 | 82.7627 | 4.2589 | 0.7982 | 0.142 | 0.1314 | 0.0361 | 0.0282 | 0.0025 | 0.0618 | 0.05
2023-03-07 10:15 | cmol_mol-1 | 0 | 10.5291 | 1.201 | 82.7508 | 4.2655 | 0.7987 | 0.1429 | 0.1334 | 0.0358 | 0.028 | 0.0026 | 0.0622 | 0.05
2023-03-07 10:30 | cmol_mol-1 | 0 | 10.5397 | 1.1853 | 82.7991 | 4.2357 | 0.7883 | 0.1407 | 0.1332 | 0.0358 | 0.0278 | 0.0024 | 0.062 | 0.05
2023-03-07 10:45 | cmol_mol-1 | 0 | 10.5533 | 1.17 | 82.8451 | 4.206 | 0.778 | 0.1386 | 0.1305 | 0.0361 | 0.028 | 0.0024 | 0.0619 | 0.05
2023-03-07 11:00 | cmol_mol-1 | 0 | 10.5662 | 1.1542 | 82.8868 | 4.1783 | 0.7701 | 0.1372 | 0.128 | 0.0358 | 0.0283 | 0.0025 | 0.0625 | 0.05
2023-03-07 11:15 | cmol_mol-1 | 0 | 10.5703 | 1.1384 | 82.9181 | 4.1656 | 0.7659 | 0.1366 | 0.1265 | 0.0355 | 0.0279 | 0.0025 | 0.0625 | 0.05
2023-03-07 11:30 | cmol_mol-1 | 0 | 10.6049 | 1.1201 | 82.9227 | 4.1531 | 0.7607 | 0.1355 | 0.1255 | 0.0351 | 0.0275 | 0.0023 | 0.0628 | 0.05
2023-03-07 11:45 | cmol_mol-1 | 0 | 10.588 | 1.1169 | 82.9595 | 4.1423 | 0.7575 | 0.1335 | 0.1241 | 0.0346 | 0.0274 | 0.0025 | 0.0636 | 0.05
2023-03-07 12:00 | cmol_mol-1 | 0 | 10.6039 | 1.1134 | 82.9123 | 4.1651 | 0.767 | 0.1349 | 0.1254 | 0.0344 | 0.0269 | 0.0021 | 0.0644 | 0.05
2023-03-07 12:15 | cmol_mol-1 | 0 | 10.626 | 1.1014 | 82.8662 | 4.1948 | 0.7727 | 0.1363 | 0.1262 | 0.0337 | 0.026 | 0.0022 | 0.0647 | 0.05
2023-03-07 12:30 | cmol_mol-1 | 0 | 10.6551 | 1.0979 | 82.8324 | 4.196 | 0.7779 | 0.1385 | 0.127 | 0.0332 | 0.0252 | 0.0023 | 0.0646 | 0.05
2023-03-07 12:45 | cmol_mol-1 | 0 | 10.6256 | 1.0989 | 82.886 | 4.1712 | 0.7745 | 0.1385 | 0.1289 | 0.0344 | 0.0263 | 0.0024 | 0.0632 | 0.05
2023-03-07 13:00 | cmol_mol-1 | 0 | 10.6391 | 1.0871 | 82.9253 | 4.1397 | 0.7658 | 0.137 | 0.1286 | 0.0349 | 0.0271 | 0.0024 | 0.063 | 0.05
2023-03-07 13:15 | cmol_mol-1 | 0 | 10.6677 | 1.0734 | 82.9256 | 4.1294 | 0.7626 | 0.1368 | 0.1276 | 0.0348 | 0.0269 | 0.0023 | 0.063 | 0.05
2023-03-07 13:30 | cmol_mol-1 | 0 | 10.6645 | 1.0714 | 82.9358 | 4.1225 | 0.7631 | 0.1368 | 0.1281 | 0.035 | 0.027 | 0.0024 | 0.0635 | 0.05
2023-03-07 13:45 | cmol_mol-1 | 0 | 10.6775 | 1.0193 | 83.1293 | 4.0112 | 0.7319 | 0.1313 | 0.1241 | 0.034 | 0.0262 | 0.0023 | 0.0629 | 0.05
2023-03-07 14:00 | cmol_mol-1 | 0 | 10.6661 | 0.9949 | 83.2548 | 3.9496 | 0.7128 | 0.1276 | 0.1213 | 0.0332 | 0.0257 | 0.0023 | 0.0618 | 0.05
2023-03-07 14:15 | cmol_mol-1 | 0 | 10.6735 | 1.0175 | 83.1636 | 3.9944 | 0.7248 | 0.1306 | 0.1222 | 0.0332 | 0.0257 | 0.0023 | 0.0622 | 0.05
2023-03-07 14:30 | cmol_mol-1 | 0 | 10.6649 | 1.0339 | 83.1064 | 4.0295 | 0.7364 | 0.1335 | 0.1239 | 0.0337 | 0.0258 | 0.0023 | 0.0597 | 0.05
2023-03-07 14:45 | cmol_mol-1 | 0 | 10.6655 | 1.0327 | 83.0924 | 4.0318 | 0.7401 | 0.1368 | 0.1275 | 0.035 | 0.0264 | 0.0025 | 0.0594 | 0.05
2023-03-07 15:00 | cmol_mol-1 | 0 | 10.6637 | 1.0032 | 83.2219 | 3.97 | 0.7145 | 0.1318 | 0.1204 | 0.0354 | 0.027 | 0.0025 | 0.0596 | 0.05
2023-03-07 15:15 | cmol_mol-1 | 0 | 10.6798 | 0.9987 | 83.2698 | 3.9351 | 0.7015 | 0.1293 | 0.1153 | 0.0328 | 0.0254 | 0.0025 | 0.0599 | 0.05
2023-03-07 15:30 | cmol_mol-1 | 0 | 10.6534 | 1.0091 | 83.2802 | 3.9396 | 0.7022 | 0.13 | 0.115 | 0.0322 | 0.0245 | 0.0024 | 0.0613 | 0.05
2023-03-07 15:45 | cmol_mol-1 | 0 | 10.6382 | 1.0206 | 83.2377 | 3.9732 | 0.7115 | 0.1322 | 0.1157 | 0.0322 | 0.0242 | 0.0025 | 0.0619 | 0.05
2023-03-07 16:00 | cmol_mol-1 | 0 | 10.6498 | 1.0294 | 83.1769 | 4.0004 | 0.7206 | 0.1341 | 0.1174 | 0.0325 | 0.0244 | 0.0024 | 0.0622 | 0.05
2023-03-07 16:15 | cmol_mol-1 | 0 | 10.6831 | 1.0348 | 83.1156 | 4.0165 | 0.7254 | 0.1342 | 0.1182 | 0.0329 | 0.0247 | 0.0023 | 0.0622 | 0.05
2023-03-07 16:30 | cmol_mol-1 | 0 | 10.6671 | 1.0455 | 83.1319 | 4.01 | 0.723 | 0.1324 | 0.1179 | 0.0328 | 0.0248 | 0.0026 | 0.062 | 0.05
2023-03-07 16:45 | cmol_mol-1 | 0 | 10.6491 | 1.0399 | 83.2216 | 3.9624 | 0.7105 | 0.1291 | 0.1161 | 0.0322 | 0.0245 | 0.0026 | 0.062 | 0.05
2023-03-07 17:00 | cmol_mol-1 | 0 | 10.6364 | 1.0433 | 83.2322 | 3.9599 | 0.7119 | 0.1287 | 0.1162 | 0.032 | 0.0242 | 0.0028 | 0.0625 | 0.05
2023-03-07 17:15 | cmol_mol-1 | 0 | 10.6913 | 1.0559 | 83.0671 | 4.0241 | 0.7343 | 0.133 | 0.1205 | 0.0323 | 0.0247 | 0.0026 | 0.0641 | 0.05
2023-03-07 17:30 | cmol_mol-1 | 0 | 10.6992 | 1.0804 | 82.9146 | 4.1021 | 0.7635 | 0.1373 | 0.1262 | 0.0334 | 0.0255 | 0.0025 | 0.0652 | 0.05
2023-03-07 17:45 | cmol_mol-1 | 0 | 10.6932 | 1.0784 | 82.9619 | 4.0778 | 0.7547 | 0.1346 | 0.1257 | 0.0332 | 0.0253 | 0.0026 | 0.0626 | 0.05
2023-03-07 18:00 | cmol_mol-1 | 0 | 10.6334 | 1.1058 | 83.0481 | 4.0453 | 0.7386 | 0.1332 | 0.123 | 0.0331 | 0.0254 | 0.0026 | 0.0615 | 0.05
2023-03-07 18:15 | cmol_mol-1 | 0 | 10.6123 | 1.1125 | 83.0695 | 4.0383 | 0.738 | 0.1324 | 0.1237 | 0.0333 | 0.0256 | 0.0024 | 0.0618 | 0.05
2023-03-07 18:30 | cmol_mol-1 | 0 | 10.5748 | 1.1269 | 83.1445 | 3.9863 | 0.7324 | 0.1328 | 0.1252 | 0.0341 | 0.0266 | 0.0026 | 0.0639 | 0.05
2023-03-07 18:45 | cmol_mol-1 | 0 | 10.5343 | 1.1488 | 83.1299 | 4.0129 | 0.7377 | 0.1345 | 0.1244 | 0.0343 | 0.0264 | 0.0028 | 0.064 | 0.05
2023-03-07 19:00 | cmol_mol-1 | 0 | 10.565 | 1.1716 | 82.9527 | 4.1232 | 0.7542 | 0.1388 | 0.1208 | 0.0334 | 0.0252 | 0.0027 | 0.0623 | 0.05
2023-03-07 19:15 | cmol_mol-1 | 0 | 10.4869 | 1.1733 | 82.939 | 4.201 | 0.7667 | 0.1423 | 0.1193 | 0.0331 | 0.0242 | 0.0027 | 0.0615 | 0.05
2023-03-07 19:30 | cmol_mol-1 | 0 | 10.5639 | 1.1594 | 82.9352 | 4.1497 | 0.7586 | 0.1422 | 0.1187 | 0.0331 | 0.0242 | 0.0029 | 0.0622 | 0.05
2023-03-07 19:45 | cmol_mol-1 | 0 | 10.5646 | 1.1393 | 82.9693 | 4.1323 | 0.7597 | 0.1434 | 0.1187 | 0.0333 | 0.024 | 0.0028 | 0.0627 | 0.05
2023-03-07 20:00 | cmol_mol-1 | 0 | 10.5818 | 1.1376 | 82.97 | 4.1183 | 0.7571 | 0.1433 | 0.119 | 0.0328 | 0.0242 | 0.0027 | 0.0633 | 0.05
2023-03-07 20:15 | cmol_mol-1 | 0 | 10.5569 | 1.151 | 83.0022 | 4.1039 | 0.751 | 0.1418 | 0.1191 | 0.0333 | 0.0248 | 0.003 | 0.063 | 0.05
2023-03-07 20:30 | cmol_mol-1 | 0 | 10.5301 | 1.1431 | 83.0149 | 4.1244 | 0.7552 | 0.1417 | 0.1182 | 0.0327 | 0.024 | 0.003 | 0.0628 | 0.05
2023-03-07 20:45 | cmol_mol-1 | 0 | 10.5355 | 1.1802 | 82.8679 | 4.1939 | 0.7789 | 0.1439 | 0.1245 | 0.0346 | 0.0257 | 0.0028 | 0.0621 | 0.05
2023-03-07 21:00 | cmol_mol-1 | 0 | 10.5502 | 1.174 | 82.8955 | 4.1698 | 0.7703 | 0.1411 | 0.1241 | 0.0343 | 0.0259 | 0.0028 | 0.062 | 0.05
2023-03-07 21:15 | cmol_mol-1 | 0 | 10.6786 | 1.1165 | 82.7767 | 4.1891 | 0.7882 | 0.1449 | 0.1294 | 0.0359 | 0.0262 | 0.0026 | 0.0618 | 0.05
2023-03-07 21:30 | cmol_mol-1 | 0 | 10.6555 | 1.1216 | 82.7209 | 4.2567 | 0.796 | 0.1456 | 0.1279 | 0.0355 | 0.0256 | 0.0023 | 0.0623 | 0.05
2023-03-07 21:45 | cmol_mol-1 | 0 | 10.7441 | 1.1071 | 82.4447 | 4.458 | 0.8132 | 0.1506 | 0.1184 | 0.0324 | 0.0223 | 0.0017 | 0.0577 | 0.05
2023-03-07 22:00 | cmol_mol-1 | 0 | 10.711 | 1.0844 | 82.6371 | 4.3644 | 0.7843 | 0.1415 | 0.1168 | 0.0317 | 0.0227 | 0.0013 | 0.0549 | 0.05
2023-03-07 22:15 | cmol_mol-1 | 0 | 10.7423 | 0.993 | 83.0004 | 4.1545 | 0.7239 | 0.1281 | 0.1094 | 0.0263 | 0.0213 | 0 | 0.051 | 0.05
2023-03-07 22:30 | cmol_mol-1 | 0 | 11.0095 | 1.0131 | 82.5754 | 4.2332 | 0.7594 | 0.1417 | 0.1114 | 0.0306 | 0.0212 | 0 | 0.0545 | 0.05
2023-03-07 22:45 | cmol_mol-1 | 0 | 11.003 | 1.0144 | 82.563 | 4.2436 | 0.7635 | 0.1442 | 0.1111 | 0.0307 | 0.0209 | 0 | 0.0556 | 0.05
2023-03-07 23:00 | cmol_mol-1 | 0 | 10.982 | 1.0251 | 82.4755 | 4.3139 | 0.7836 | 0.148 | 0.1132 | 0.0314 | 0.021 | 0 | 0.0563 | 0.05
2023-03-07 23:15 | cmol_mol-1 | 0 | 10.8158 | 1.0523 | 82.5312 | 4.3641 | 0.8023 | 0.1501 | 0.1181 | 0.033 | 0.0229 | 0.0015 | 0.0589 | 0.05
2023-03-07 23:30 | cmol_mol-1 | 0 | 10.657 | 1.1004 | 82.8038 | 4.2242 | 0.7772 | 0.1425 | 0.1223 | 0.0336 | 0.0249 | 0.0023 | 0.0617 | 0.05
2023-03-07 23:45 | cmol_mol-1 | 0 | 10.7142 | 1.0774 | 82.7589 | 4.2256 | 0.7817 | 0.1438 | 0.1238 | 0.034 | 0.0255 | 0.0022 | 0.0631 | 0.05
2023-03-08 | cmol_mol-1 | 0 | 10.7123 | 1.0856 | 82.8079 | 4.1749 | 0.7743 | 0.1442 | 0.1239 | 0.0342 | 0.0257 | 0.0026 | 0.0644 | 0.05

## USM_unc
Input variable | Input value | Uncertainty | Unit | Confidence interval | Distribution
Min_velocity | 0.3 |  | ms-1
Max_velocity | 20 |  | ms-1
Calibration reference |  | 0.2 | % | 95% | Normal
Calibration repeatability |  | 0.1 | % | 95% | Normal
Calibration deviation |  | 0.3 | % | 95% | Normal
Field uncertainty |  | 0.1 | % | 95% | Normal
Installation effects |  | 0 | % | 95% | Normal
Low_limit | 500 |  | quantity_m3
High_limit | 500000 |  | quantity_m3

## T_unc
Input variable | Input value | Uncertainty | Unit | Confidence interval | Distribution
Time_between_calibrations | 12 |  | months
Ambient_Temp_at_calibration | 20 |  | C
Temp_transmitter_element |  | 0.1 | C | 99% | Normal
Stability_transmitter |  | 0.1 | %/24mo | 99% | Normal
RFI_effects |  | 0.1 | C | 99% | Normal
Ambient_Temp_effect |  | 0.0015 | C/C | 99% | Normal
Stability_temp_element |  | 0.05 | C | 95% | Normal
Misc |  | 0.1 | C | 95% | Normal
Low_limit | -40 |  | C
High_limit | 200 |  | C

## P_unc
Input variable | Input value | Uncertainty | Unit | Confidence interval | Distribution
Max_calibrated_static_pressure | 120 |  | barg
Min_calibrated_static_pressure | 50 |  | barg
Upper_range_limit | 138 |  | barg
Time_between_calibrations | 12 |  | months
Ambient_temp_at_calibation | 20 |  | C
Transmitter |  | 0.05 | %span | 99% | Normal
Stability |  | 0.1 | %URL/yr | 95% | Normal
RFI_effects |  | 0.1 | %spen | 99% | Normal
Ambient_temp_effect |  | 0.03 | %span/28C | 99% | Normal
Atm_pressure |  | 0.09 | barg | 95% | Normal
Misc |  | 0.1 | barg | 95% | Normal
Low_limit | 2.9995 |  | MPa
High_limit | 100 |  | MPa
Note on D7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: allowing the user to define the uncertainties in many different ways may be challenging to handle on the backend

## PT_unc (hidden)
Input variable | Uncertainty | Unit | k | Distribution
T | 0.3 | C | 2 | Normal
P | 0.2 | bara | 2 | Normal

## Composition_unc
Comp | H2 | N2 | CO2 | C1 | C2 | C3 | iC4 | nC4 | neoC5 | iC5 | nC5 | C6 | He
Uncertainty | 0.04 | 0.004 | 0.002 | 0.172 | 0.018 | 0.009 | 0.0044 | 0.0036 | 0.001 | 0.0018 | 0.0013 | 0.0008 | 0.004
Unit | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1 | cmol_mol-1
Confidence interval | 95%
Distribution | Normal

## lists
Quality meters | Configuration | Quality measurement | EoS
USM | Single | Sampling | GERG-2008
Coriolis | Dual in parallel | OnlineGC | AGA-8
| Dual in series | Calorimeter
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Quantity?